The serial communication protocol

- Definition

The communication protocol between computer and EDFA module has been defined in this document. In this protocol the computer act as MASTER, while the EDFA module as SLAVE. DOWN LINK refers to the computer send data to the EDFA module and the reverse transmission is named UP LINK.

ALL the data in this document are Hexadecimal. For example, 0xAA and AA are same meaning.

- Serial port settings

Baud rates: 9600bits/s

Data bits: 8

Stop bit: 1

Parity bit: 0

- Frame format

ALL communication apply MASTER SLAVE response mode. i.e. the MASTER send down a string of data which include a command byte. The SLAVE send up a string of corresponding data while it received the command from the HOST.

The data format of DOWN LINK

| Header | Command Byte | Length of Data | Data Bytes | Parity |
| --- | --- | --- | --- | --- |
| 2 bytes | 1 byte | 1 byte | n bytes | 2 bytes |

Which Header include 2 bytes. The first is AA and the second is 55.

The description of Command Byte:

0x11	Pump ON

0x33	Pump OFF

0x22	Get the information of EDFA

0x44	Set the pump power

0x99	Save current pump power as default

Length of Data is the number of Data Bytes. There will have no Data Bytes in the frame if it has no data to be sent.

Every data to be sent to EDFA module should be separated into a lower byte and a high byte. The lower byte be sent in front of the high byte.

Parity is the sum of Data Bytes. It should be separated into a lower byte and a high byte. The lower byte be sent in front of the high byte.

The data format of UP LINK

| Header | Command Byte | Length of Data | Data Bytes | Parity |
| --- | --- | --- | --- | --- |
| 2 bytes | 1 byte | 1 byte | n bytes | 2 bytes |

Which

Header include 2 bytes. The first is 55 and the second is AA.

The other bytes are the same with DOWN LINK frame.

- Examples

DOWN LINK

- cmd=0x44：Set the output power

| AA | 55 | 44 | 02 | (pump power%256) The lower byte | (pump power/256) The high byte | The lower byte of parity | The high byte of parity |
| --- | --- | --- | --- | --- | --- | --- | --- |

which：the unit of pump power is mW

- cmd=0x99：set current power as default

| AA | 55 | 99 | 00 | 00 | 00 |
| --- | --- | --- | --- | --- | --- |

- cmd=0x22: get information from EDFA module

| AA | 55 | 22 | 00 | 00 | 00 |
| --- | --- | --- | --- | --- | --- |

- cmd=0x11: Pump ON

| AA | 55 | 11 | 00 | 00 | 00 |
| --- | --- | --- | --- | --- | --- |

- cmd=0x33: Pump OFF

| AA | 55 | 33 | 00 | 00 | 00 |
| --- | --- | --- | --- | --- | --- |

UP LINK

| 1 | 55 |
| --- | --- |
| 2 | AA |
| 3 | cmd |
| 4 | 0X1A |
| 5 | The lower byte of output power |
| 6 | The high byte of output power |
| 7 | Reserved |
| 8 | Reserved |
| 9 | Reserved |
| 10 | Reserved |
| 11 | Reserved |
| 12 | Reserved |
| 13 | Reserved |
| 14 | Reserved |
| 15 | Reserved |
| 16 | Reserved |
| 17 | The lower byte of pump temperature |
| 18 | The high byte of pump temperature |
| 19 | Reserved |
| 20 | Reserved |
| 21 | The lower byte of pump current |
| 22 | The high byte of pump current |
| 23 | Reserved |
| 24 | Reserved |
| 25 | The lower byte of pump voltage |
| 26 | The high byte of pump voltage |
| 27 | Reserved |
| 28 | Reserved |
| 29 | Reserved |
| 30 | Reserved |
| 31 | The lower byte of parity |
| 32 | The high byte of parity |

which：

1. the current of pump = (byte22*256+byte21)  (mA)

2. the temperature of pump =(byte18*256+byte17) / 100  (℃)

3. the voltage of pump =(byte26*256+byte25) / 1000  (V)

4. the output power =(byte6*256+byte5)/10  (mW)

- Port Definition

| DB25_1 | +5VDC | DB25_14 | GND |
| --- | --- | --- | --- |
| DB25_2 | +5VDC | DB25_15 | GND |
| DB25_3 | +5VDC | DB25_16 | GND |
| DB25_4 | +5VDC | DB25_17 | GND |
| DB25_5 | Reserved | DB25_18 | UART-TXD |
| DB25_6 | Reserved | DB25_19 | Reserved |
| DB25_7 | Reserved | DB25_20 | Reserved |
| DB25_8 | Reserved | DB25_21 | Reserved |
| DB25_9 | Reserved | DB25_22 | UART-RXD |
| DB25_10 | Reserved | DB25_23 | UART-GND |
| DB25_11 | Reserved | DB25_24 | Reserved |
| DB25_12 | Reserved | DB25_25 | Reserved |
| DB25_13 | Reserved |  |  |